Evaluation of marginal fit of IPS e.max ZirCAD crowns fabricated by a CAD-CAM CEREC inLAB MC X5 system and LAVA milling center

ABSTRACT

Background. Marginal fit is critical for the success and longevity of a dental restoration. Zirconia crowns can be fabricated either chair-side, in a dental laboratory or in a milling center, with differing marginal fits produced by each method. However, discussion of the marginal fit of zirconia crowns when different fabrication methods are compared is lacking in the literature. Purpose. This study aimed to determine the existence of a significant difference between the marginal discrepancy (MD) and absolute marginal discrepancy (AMD) of computer‐aided design, and computer‐aided machining (CAD-CAM) used in a dental laboratory and a milling center for producing monolithic zirconia crowns.

Methods. The marginal fit of 30 zirconia crowns cemented to typodont teeth was evaluated by means of a sectioning technique. Fifteen crowns were fabricated with a CEREC inLAB MC X5 from IPS e.max ZirCAD blocks. Fifteen crowns were fabricated using a LAVA milling center from LAVA Plus Zirconia Blocks. The 30 crowns were sectioned with a precision saw and MD and AMD were subsequently measured using a light microscope. Data were analyzed using the one-way ANOVA technique to investigate significant differences in the marginal fit between the two fabrication systems (α=.05).

Results. Mean AMD dimensions for zirconia crowns fabricated by the CEREC inLAB were 85 µm, and for the LAVA milling center 133 µm. The MD dimensions were 53 µm for the CEREC inLAB and 61 µm for the LAVA milling center. The AMD of the CEREC inLAB system was significantly smaller (P<.05). There was no significant difference between the two systems regarding the MD dimensions.

Conclusions. The CEREC inLAB system demonstrated significantly better marginal fit in relation to the AMD. However, no difference between the systems was found in the MD. Monolithic zirconia crowns fabricated by the CAD-CAM CEREC inLAB system and the LAVA system milling center were within the clinically acceptable range for marginal fit.

KEYWORDS

# CAD-CAM, zirconia, crown, marginal fit, CEREC, LAVA

INTRODUCTION

In recent decades, increasing demand from patients for natural-appearing dental restorations has led to the development of all-ceramic materials with improved mechanical characteristics that ensure suitable longevity. These are now replacing traditional metal-ceramic restorations [1,2,3]. The introduction of CAD-CAM technology allows for the use of materials such as zirconia, which are free of metal, in dental restorations [1].

Zirconia is a polycrystalline ceramic without a glassy phase and exists in several temperature-dependent forms. At room temperature it exists in a monoclinic crystalline form, changing to a tetragonal and cubic crystalline form when sintered [4]. The cooling from cubic to tetragonal results in an expansion of 2.3 % and from tetragonal to monoclinic of 4.2 %. These expansions are the cause of cracks and hence there is a need to stabilize the tetragonal form. The most common method of stabilizing the tetragonal phase and maintaining zirconia in a metastable condition at room temperature is achieved via the addition of a small amount of yttria to the zirconia [3,4,5]. Such treatment produces a stronger material than other available ceramics. The zirconia is a biocompatible material with high mechanical properties of 1200 HV hardness, 900–1200 MPa flexural strength and fracture toughness of 6–8 MPa m1/2 [6,7].

Zirconia restorations fabricated by CAD-CAM technology can be produced chair-side, in a laboratory or in a milling center. The restorations are processed either by soft machining of pre-sintered blanks with enlarged contours followed by sintering at high temperature during which they shrink to their desired and final size, or by hard machining of fully sintered blocks [8].

Superior marginal fit is an important characteristic for the success and longevity of dental restorations. Poor marginal fit results in plaque retention and microleakage; this can lead to secondary dental caries, pulpal lesions, periodontal disease, and bone loss [9,10]. Dental literature includes significant investigation of the accuracy of marginal fit. Marginal discrepancies of between 50 and 120 μm are considered clinically acceptable as regards longevity of the restoration, while more restrictive studies proposed marginal discrepancies of less than 100 µm [11,12]. An in vivo study of more than 1000 crowns found a greater association between a marginal discrepancy of less than 120 µm and higher longevity [13]. Studies on the marginal fit of zirconia copings fabricated by CAD-CAM have reported measured marginal discrepancies of as low as 10 μm and as high as 110 μm, with most being less than 80 μm [14]. In contrast, studies of full zirconia crown marginal fit are lacking.

The definition of marginal fit can differ and depends on the gap measured in the studies. Holmes et al. defined the marginal discrepancy (MD) as the perpendicular measurement from the cervical margin of the restoration to the preparation margin, while the AMD is measured from the cervical margin of the restoration to the cavosurface of the preparation [15]. The MD represents the surface of the cement that is exposed to the oral environment and can be dissolved, resulting in microleakage. The AMD is indicative of the under- or over-extension of the restoration margins relative to the margins of the preparation and plays a significant role in plaque accumulation (Fig. 1) [15].

Studies that have compared the marginal fit of zirconia copings to other ceramic restorations show higher accuracy for zirconia [16]. The marginal fit of zirconia copings produced using different CAD-CAM system has also been investigated [17,18]. However, since studies of monolithic zirconia crowns produced using different fabrication methods are lacking and the use of monolithic zirconia crowns is increasing, a comparison of fabrication methods is justified. The purpose of this in vitro study was therefore to compare the marginal fit of monolithic zirconia crowns produced using two fabrication methods: dental laboratory and milling center. The null hypothesis was that no difference would be found in the marginal fit of the fabrication methods.

Many methods exist for the evaluation of marginal fit [12,16,17,19,20]. In this study, the sectioning method was used to investigate the two parameters of marginal fit, AMD and MD [17,20].

MATERIAL AND METHODS

The following method is the same as that previously published by Dolev et al. and will be described here only briefly [21]. Mandibular left first molar, typodont teeth (FLUX 8634; Columbia Dentoform) were used as abutment. For the CEREC inLAB system group, 15 crowns were prepared in the dental laboratory (TOTALI - AMIR LIFF LTD, Tel Aviv, Israel) from partially sintered zirconia blocks (IPS e.max ZirCAD; Ivoclar Vivadent) using a milling unit (CEREC inLAB MC X5; Dentsply Sirona), followed by sintering (Ceramill Therm 1; Amann Girrbach) to produce completely sintered crowns (Fig. 2). Fifteen typodont teeth were sent to a milling center (LAVA; 3M ESPE) for scanning (Lava Scan ST; 3M ESPE) and design (LAVA Design 5; 3M ESPE). The zirconia crowns were fabricated by partially sintered zirconia blocks (LAVA Plus Zirconia Blocks; 3M ESPE) using a CAD milling machine (LAVA Form; 3M ESPE) followed by sintering (LAVA Furnace 200; 3M ESPE) for production of the final crowns.

The crowns were cemented with self-adhesive resin cement (Rely X U-200; 3M ESPE) and then sectioned with a cutting machine (Izomet Plus precision saw; Buehler), creating four specimens from each crown: Mesio-Buccal (MB), Disto-Buccal (DB), Disto-Lingual (DL), and Mesio-Lingual (ML). In each specimen, the AMD and MD were measured in two locations (Fig. 3) using a light microscope (Axioplan 2; Zeiss) at ×110 magnification [21].

Repeat measurements using the one-way ANOVA statistical test were carried out (α=.05) to examine significant differences between the groups.

RESULTS

The mean values with standard errors for the AMD and MD dimensions for crowns fabricated by the CEREC inLAB and LAVA milling center are shown in Figures 4 and 5. The CEREC inLAB presented smaller AMD values than the LAVA milling center (Fig. 4). The MD values of CEREC inLAB crowns were smaller than those produced in the LAVA milling center except for Mid-L/ML, Mid-L/DL, and Mid-M/ML locations (Fig. 5).

The CEREC inLAB generated an overall mean ± standard error (SE) value for AMD of 85 ±2 µm, while the LAVA milling center generated an overall mean ±SE value for AMD of 133 ±4 µm. The overall mean ±SE MD value was 53 ±2 µm for crowns milled by CEREC inLAB and 61 ±3 µm for the crowns milled using the LAVA milling center technique. The statistical outcome showed significant differences for AMD (df=1, F=35.081; P=.000) whereas MD yielded no significant differences (df=1, F=1.799; P=.191) between CEREC inLAB and the LAVA milling center. The MD 95% confidence intervals found a lower bound of 43 µm for CEREC inLAB and 52 µm for the LAVA milling center, and an upper bound of 62 µm for CEREC inLAB and 71 µm for the LAVA milling center.

# DISCUSSION

In accordance with the study results, the null hypothesis regarding the AMD parameter was rejected, with the dental laboratory using the CEREC inLAB CAD-CAM displaying a significantly lower gap (85 ±2 µm) compared to the LAVA milling center (133 ±4 µm). For the MD parameter, the null hypothesis was not rejected since no statistically significant differences were found between the 2 systems.

Beuer et al. identified significant differences when examining the MD of three-unit zirconia frameworks [17]. They found a smaller mean MD value in the milling center (29.1 µm) than in frameworks produced by CEREC inLAB (56.6 µm). However, this study made use of an Etkon milling center. Rajan et al. compared the marginal fit of zirconia coping produced by CEREC inLAB with that of the CERAMILL system and found significant differences, whereby CEREC inLAB copings had better adaptation than CERAMILL [18]. Marginal fit for CERAMILL was 83 μm and for the CEREC InLAB MC XL was 68 μm [18].

In their systemic review of the fit of zirconia restorations, Abduo et al. [22]. indicated the difficulty in comparing the many studies existing on the marginal gap of zirconia given the different methodology used in each study [22], including the sectioning technique [17,20], use of microcomputed tomography [16,19], and silicone paste technique [12]. Additionally, each study examined and compared different parameters of marginal fit, MD and AMD being a few of the many that were described by Holmes et al. [15]. For this reason, there is a need for standardization.

Crown cementation has been used in the present study to reproduce the clinical conditions of the crown-abutment relationship. Earlier publications have found that crown cementation has a negative effect on the marginal fit, which increases after cementation [19,23].

McLean and Fraunhofer showed that crown marginal discrepancies ranging up to 120 μm were clinically acceptable [13]. According to the 95% confidence interval, the present study yielded MD values within the clinically acceptable range.

This study examined MD and AMD in 4 surfaces, buccal, lingual, mesial, and distal. This was done without comparing parameters between the different surfaces on the grounds that this is clinically irrelevant since this study used model teeth with a constant finish line, in an in-vitro setting.

Several limitations were identified in the study. The study was conducted in vitro with typodont teeth used as abutments instead of natural teeth, and finger pressure was used to lute the crowns. These characteristics differ from those of the intraoral environment. The cemented crowns were cut with a disk, which can have negative effect on the quality of specimens and the reading of marginal fit. Additionally, the cement thickness in the occlusal area that affects the internal marginal fit and affects the seating quality of the crowns was not measured, and could therefore have influenced the MD. The 2 CAD-CAM systems used zirconia blocks manufactured by different companies which may also have affected the results.

The study revealed that when using a well-known, established CAD-CAM system, zirconia monolithic crowns are good treatment options as regards marginal fit during tooth restoration, which is of relevance given the popularity of zirconia monolithic crowns as a treatment option. Because these systems are constantly developing with the arrival of new manufacturers, further in vitro and in vivo studies are needed to substantiate these results.

CONCLUSIONS

Within the limitations of this study, the following conclusions were drawn:

1. The CEREC inLAB system shows a significantly smaller AMD than the LAVA milling center.

2. No significant difference was found in MD between the systems.

3. Monolithic zirconia crowns fabricated by the CEREC inLAB system and the LAVA milling center produced MD values within the clinically acceptable standard (120 μm).

4. There is a need for standard rules and guidance when comparing marginal fit between different CAD-CAM systems.

NOTES

Avi Meirowitz and Yoli Bitterman contributed equally to this work.

List of abbreviations

CAD-CAM—computer‐aided design, and computer‐aided machining

MD—Marginal discrepancy

AMD—Absolute marginal discrepancy

SE—standard error

MB—Mesio Buccal

DB—Disto Buccal

DL—Disto Lingual

ML—Mesio Lingual

Declarations

Ethics approval and consent to participate:

Not applicable

Consent for publication:

Not applicable

Availability of data and material:

The datasets used and/or analysed during the current study are available from the corresponding author on reasonable request.

Competing interests:

The authors declare that they have no competing interests

Funding:

The authors have no funding to report

Authors' contributions:

All the authors made substantial contributions to the present study. AM designed the study, and reviewed and edited the manuscript; YB analyzed and interpreted the results and was a major contributor in writing the manuscript; SL performed the experiment; EM performed critical revision of the paper; ED was responsible for conceptualization, methodology, supervision, writing (review and editing). All authors read and approved the final manuscript.

REFERENCES

1. Raigrodski AJ, Hillstead MB, Meng GK, Chung KH. Survival and complications of zirconia-based fixed dental prosthesis: a systematic review. J Prosthet Dent 2012; 107:170-177. Review.

2. Biscaro L, Bonfiglioli R, Soattin M, Vigolo P. An in vivo evaluation of fit of zirconium-oxide based ceramic single crowns, generated with two CAD/CAM systems, in comparison to metal ceramic single crowns. J Prosthodont 2013; 22:36-41.

3. Zarone F., Russo S., Sorrentino R. From porcelain-fused-to-metal to zirconia: Clinical and experimental considerations. Dent. Mater. 2011;27:83–96.

4. I Denry, JR Kelly. State of the art of zirconia for dental applications. Dent Mater 2008;24: 299-307

5. Christel P, Meunier A, Heller M, Torre JP, Peille CN. Mechanical properties and short-term in-vivo evaluation of yttrium-oxide-partially-stabilized zirconia. J Biomed Mater Res 1989;23:45-61.

6. Piconi C, Maccauro G. Zirconia as a ceramic biomaterial. Biomaterials, 1999;20:1-25.

7. T. Kosmac, C. Oblak, P. Jevnikar, N. Funduk, L. Marion. The effect of surface grinding and sandblasting on flexural strength and reliability of Y-TZP zirconia ceramic. Dent Mater 1999; 15:426-433.

8. F. Filser, P. Kocher, L.J. Gauckler. Net-shaping of ceramic components by direct ceramic machining. Assembly Autom 2003; 23:382-390.

9. J.A. Sorensen. A rationale for comparison of plaque-retaining properties of crown systems. J Prosthet Dent 1989; 62:264-269.

10. H.G. Kashani, S.C. Khera, I.A. Gulker. The effects of bevel angulation on marginal integrity. J Am Dent Assoc 1981; 103:882-885.

11. Suárez MJ, González de Villaumbrosia P, Pradíes G, Lozano JF. Comparison of the marginal fit of procera AllCeram crowns with two finish lines. Int J Prosthodont 2003; 16:229-32.

12. Coli P, Karlsson S. Fit of a new pressure-sintered zirconium dioxide coping. Int J Prosthodont 2004; 17:59-64.

13. McLean JW, von Fraunhofer JA. The estimation of cement film thickness by an in vivo technique. Br Dent J 1971;131:107-11.

14. P. Boitelle, B. Mawussi, L. Tapie, O. Fromentin. A systematic review of CAD/CAM fit restoration evaluations. J Oral Rehabil 2014; 41:853-874.

15. Holmes JR, Bayne SC, Holland GA, Sulik WD. Considerations in measurement of marginal fit. J Prosthet Dent 1989; 62:405-8.

16. Riccitiello F, Amato M, Leone R, Spagnuolo G, Sorrentino R. In vitro Evaluation of the Marginal Fit and Internal Adaptation of Zirconia and Lithium Disilicate Single Crowns: Micro-CT Comparison Between Different Manufacturing Procedures. Open Dent J 2018; 12:160-172.

17. Beuer F, Aggstaller H, Edelhoff D, Gernet W, Sorensen J. Marginal and internal fits of fixed dental prostheses zirconia retainers. Dent Mater J 2009; 25: 94-102.

18. Rajan BN, Jayaraman S, Kandhasamy B, Rajakumaran I. Evaluation of marginal fit and internal adaptation of zirconia copings fabricated by two CAD-CAM systems: An in vitro study. The Journal of the Indian Prosthodontic Society 2015; 15:173-178.

19. Demir N, Ozturk AN, Malkoc MA. Evaluation of the marginal fit of full ceramic crowns by the microcomputed tomography (micro-CT) technique. Eur J Dent 2014; 8:437-44.

20. Sulaiman F, Chai J, Jameson LM, Wozniak WT. A comparison of the marginal fit of In-Ceram, IPS Empress, and procera crowns. Int J Prosthodont 1997; 10:478-84.

# 21. Dolev E, Bitterman Y, Meirowitz A. Comparison of marginal fit between CAD-CAM and hot-press lithium disilicate crowns. J Prosthet Dent. 2018 Jun 29. [Epub ahead of print]

# 22. Abduo J, Lyons K. Swain M. Fit of zirconia fixed partial denture: a systematic review. J Oral Rehabil 2010;37:866-876.

23. Borges GA, Faria JS, Agarwal P, Spohr AM, Correr-Sobrinho L, Miranzi BA. In vitro marginal fit of three all-ceramic crown systems before and after cementation. Oper Dent 2012;37:641-649.

FIGURES

Figure 1. Discrepancies between crown and abutment finish line. (a) Marginal discrepancy (MD); (b) Absolute marginal discrepancy (AMD).21

Figure 2. Lower left first molar typodont tooth and corresponding CEREC inLAB zirconia crown.

Figure 3. Illustration of sectioned mandibular left first molar. Colored points demonstrate 8 measurement locations of finish line discrepancies.

Figure 4. Comparison of mean values and standard errors of AMD at different marginal area locations for CEREC inLAB system and LAVA milling center crowns.

Figure 5. Comparison of mean values and standard errors of MD at different marginal area locations for CEREC inLAB system and LAVA milling center crowns.